Response Letter

Paper:  JBHI-04130-2024

Title: Ultra-High Quality ECG Compression for IoMT Application Using Temporal Convolutional Auto-Encoder with Improved RVQ

Authors: Yeyi Guan, Jingyuan Qiao, Kailin Zhuang, Xinjie Zhang, Student Member, IEEE, Ruowen Qu, Fenghan Zhou, Jianxiu Jin and Lin Shu, Senior Member, IEEE

Response to Editorial Comments

The authors thank the handling associate editor (AE) and anonymous reviewers for their valuable and constructive comments. The suggestion and detailed comments from the AE and reviewers are greatly helpful for improving the quality of our manuscript. We have provided the point-by-point responses to each comment raised by all reviewers. The modifications have been highlighted in the revised manuscript.

Response to Review #1

- Although the author points out that CR is a commonly used evaluation metric in the field of ECG compression, from the perspective of the long-established rate-distortion theory in the field of lossy compression, CR is a non-ideal metric that can be misleading, as it lacks a strict correlation with the actual compression efficiency in the compression process. In their response, the author can only confirm that, under a specific experimental scenario, there might be an approximate relationship between the entropy rate and CR under certain conditions; however, in broader application scenarios, the deviation between the two may significantly increase. Particularly in comparative experiments, the CR metric does not provide effective guidance for evaluating the actual performance of different compression methods—for instance, it is possible that Method A has a lower CR than Method B, yet Method A has a lower entropy rate and higher actual compression efficiency.

Furthermore, the author mentions in their response the empirical entropy and the entropy rate output by the Transformer-based entropy estimator, indicating that after achieving accurate distribution estimation, the compression rate is significantly improved. This phenomenon further highlights the deviation between CR and actual compression performance. The author attributes this to "the excessively long encoding time after adding the entropy estimator," but it must be clarified: the root cause of this issue is not the entropy estimator itself, but rather the inherent complexity of the Transformer model; moreover, the encoding and decoding time of entropy models constructed by neural networks is closely related to the actual hardware equipment and specific implementation. In fact, there are various types of entropy models in the field of lossy compression with lower encoding and decoding costs. Therefore, "discarding entropy coding to improve encoding and decoding speed" is not a reasonable justification. If the author wishes to further optimize encoding and decoding efficiency, it is necessary to comprehensively consider the forward optimization of the encoding and decoding process of deep learning-based VQ models, rather than simply discarding the entropy coding module.

Of course, conducting entropy rate tests for all methods is quite challenging. Accordingly, it is recommended that the author supplement the paper with an explanation of why the field chooses to report CR instead of the entropy rate (i.e., the uniqueness compared to other compression fields), while also providing a theoretical explanation of the consistency between CR and the entropy rate in the field of ECG compression, as well as the unique advantages of VQ methods in compression tasks in this field. This will help researchers in the compression field more fully understand the effectiveness of the method proposed in this paper.

Response:

We sincerely thank the reviewer for this insightful comment grounded in classical Rate-Distortion Theory. We unequivocally acknowledge that, from an information-theoretic perspective, Entropy Rate (BPS) is the fundamental measure of compression efficiency, and CR is indeed a non-ideal proxy that can be misleading across heterogeneous methods. We fully agree that the reviewer’s hypothetical scenario—where Method A achieves a lower CR yet a lower true entropy rate than Method B—is theoretically valid and exposes CR's fundamental limitation as a universal metric.

However, after a comprehensive investigation of ECG compression literature spanning four decades, we found that the ECG domain exhibits unique characteristics that have historically made CR a pragmatically effective (though not theoretically superior) metric for intra-domain comparison. Our rationale for retaining CR in this specific study is grounded in three factors:

1. Unprecedented Data Homogeneity in Standard Benchmarks Unlike general signal compression (e.g., audio, images) which must handle vast format diversity, the canonical MIT-BIH Arrhythmia Database—upon which the vast majority of ECG compression literature is built—enforces near-identical acquisition parameters: a unified 360 Hz sampling rate, fixed 11-bit resolution, and consistent 30-minute duration. This benchmark-enforced homogeneity artificially creates a strong correlation between CR and BPS within this specific closed system, rendering CR a de facto bandwidth proxy for intra-dataset comparisons. We acknowledge this is not a theoretical triumph of the metric itself, but an artifact of historical data standardization.

2. Domain-Specific Benchmarking Legacy ECG compression research emerged from clinical telemetry systems where the compression ratio directly dictated the transmission duty cycle and battery life. Consequently, the community adopted CR because it translated linearly to power savings in that specific hardware context. While we do not claim this legacy choice was theoretically optimal, adhering to it is necessary for backward compatibility. Abandoning CR would sever quantitative continuity with widely cited baselines [1][2][3][4][5]. We report CR primarily to place our work within this substantial body of existing clinical validation data.

3. Hardware-Redundancy Isolation Beyond legacy, CR offers specific utility in distinguishing algorithmic efficiency from source entropy across heterogeneous hardware (typically 10-bit to 24-bit).

- The Limitation of BPS: BPS correctly measures the absolute transmission burden. For instance, a 24-bit recording compressed to 2.0 BPS requires more bandwidth than an 11-bit recording compressed to 1.5 BPS. However, relying solely on BPS conflates the source entropy (input difficulty) with algorithmic performance.

- The Specific Utility of CR: In the example above, the algorithm processing the 24-bit source has achieved a superior reduction ratio, effectively stripping away massive amounts of hardware-dependent redundancy (such as thermal noise in the LSBs). CR explicitly normalizes against the raw input bit-depth, making it effective at isolating this algorithmic redundancy removal from the hardware-level redundancy.

Crucially, we must transparently limit the scope of this justification. We explicitly acknowledge that CR remains a compromise metric, functioning primarily as a legacy artifact. Despite its utility for hardware normalization in this specific study, we unequivocally recognize that Entropy Rate (BPS) constitutes the theoretical gold standard for quantifying true information density. Therefore, in the revised manuscript (Section V-D), we have adopted a rigorous transparency strategy:

- We report CR primarily as a "legacy comparator" for backward compatibility. Furthermore, we explicitly contextualize it as a necessary normalization tool against source entropy (e.g. thermal noise) in high-resolution sensors, distinguishing algorithmic redundancy removal from hardware variance.

- We have added a dedicated subsection explicitly stating that CR's reliability breaks down in cross-platform comparisons. We propose a future "dual-reporting strategy (CR+BPS)" to ensure validity across differing bit-depths and outline a roadmap for exploring non-autoregressive architectures.

Furthermore, we appreciate the reviewer's suggestion to consider alternative entropy coding mechanisms and acknowledge that relying solely on the latency of Transformer-based estimators is an incomplete justification for temporarily omitting entropy coding. We recognize that efficient parallelizable models (e.g., Scale Hyperprior) exist and can effectively capture correlations with lower computational costs. However, our decision to defer entropy coding in this work is not merely a compromise for inference speed, but a deliberate design choice to avoid the architectural complexity inherent to state-dependent bitstream generation.

While entropy coding (via Arithmetic Coding or ANS) excels at compression, it generates a serialized bitstream where decoding depends on the internal state. Although strategies like block-based segmentation can theoretically restore error resilience and random access, they impose additional synchronization overhead and system complexity. Moreover, it is important to distinguish the fundamental difference in redundancy handling between Scalar Quantization (SQ) and Vector Quantization (VQ) frameworks. In SQ-based architectures (where Scale Hyperprior excels), the quantizer is typically uniform, leaving a significant burden on the entropy model to compress the resulting distribution (often modeled as Gaussian). In such cases, removing the entropy model would result in substantial performance degradation. In contrast, VQ based approach shifts the compression burden to the codebook learning phase. By optimizing the codebook utilization, the VQ network inherently 'packs' information into the indices, resulting in a distribution that is already highly compact (approaching a uniform distribution, maximizing Marginal Entropy). Therefore, our research prioritizes structural feature extraction via the VQ network to empirically maximize the Marginal Entropy of the indices directly. By ensuring each token carries high effective information density, we aim to reduce the immediate necessity for a subsequent entropy coding stage in this safety-critical medical monitoring context. This allows us to preserve intrinsic error resilience and native O(1) random access without the engineering overhead required to force these capabilities into an entropy-coded stream.

To rigorously validate this design choice, we conducted a comprehensive comparative evaluation between our native 10-bit fixed-length coding scheme (determined by the codebook size of 1024) and a spectrum of entropy coding strategies, ranging from industry-standard general-purpose compressors (zlib [6][7] and gzip [8][9]) to a dedicated Arithmetic Coder and a theoretical analysis of the Marginal Entropy (H0). All metrics were averaged across the entire patient dataset to ensure statistical robustness. To rule out implementation bias and strictly test compressibility limits, we evaluated three distinct scenarios: mapping indices to standard uint16 containers to simulate typical engineering overhead, explicitly compacting indices via bit-packing to eliminate this overhead, and benchmarking against algorithmic and theoretical lower bounds.

As detailed in Table.1-RL, Table.2-RL, and Table.3-RL, the results provide critical insights into the efficiency of our approach. As expected, mapping indices to standard containers resulted in significant compression expansion (approximately -16% to -19%) due to padding redundancy. More critically, even with optimized bit-packing to eliminate this overhead, the gains from general-purpose compressors remained negligible, peaking at only +0.94%. Most importantly, our analysis of the entropy limit revealed a narrow margin for improvement. We explicitly acknowledge that Marginal Entropy (H0) represents the theoretical lower bound under the assumption of token independence, and does not account for higher-order temporal correlations (which would be captured by Conditional Entropy). However, exploiting these temporal dependencies inherently requires state-dependent autoregressive modeling, which contradicts our requirement for stateless random access. Under this constraint, the gap between Fixed-Length and H0 is remarkably narrow, ranging from only 2.96% to 6.36%. Even a dedicated Arithmetic Coder was able to capture only a fraction of this potential, yielding a maximum gain of +2.56%.

These empirical results demonstrate that our VQ optimization has effectively maximized the Marginal Entropy, achieving a high-density representation where statistical redundancy is minimal. In light of the fact that the maximum achievable gain via complex entropy coding is limited to approximately 2.6%, we suggest that this modest improvement not fully outweigh the introduction of architectural complexity—particularly regarding the preservation of intrinsic error resilience and native O(1) random access.

Looking forward, while fixed-length coding ensures the robustness essential for current deployments, we recognize that capturing residual high-order temporal dependencies (minimizing Conditional Entropy) remains a valuable frontier. Although standard autoregressive models currently impose unjustifiable architectural complexity for our specific constraints, we see potential in lightweight, block-wise, or non-autoregressive architectures. Our future work will explore these directions to harness residual dependencies without compromising the critical real-time and robustness capabilities established in this work.

Following the insightful suggestion, we have revised Section V-D (Detailed Compression Metrics and Justification) in the manuscript to explicitly clarify the rationale for prioritizing CR in the ECG domain, mathematically demonstrate the consistency between CR and Bit Rate, and highlight the structural advantages of VQ methods. The corresponding content can be found in Section V-D and Table VII, on page 12 and 13.

Table.1-RL Efficiency Comparison: Fixed-Length vs. Zlib

| CR | Fixed | Zlib (standard) |  | Zlib (bitpacked) |  |
| --- | --- | --- | --- | --- | --- |
|  | BPS | BPS | Gain | BPS | Gain |
| 88 | 0.1250 | 0.1458 | -16.60% | 0.1239 | +0.94% |
| 58 | 0.1876 | 0.2211 | -17.90% | 0.1874 | +0.11% |
| 44 | 0.2501 | 0.2975 | -18.96% | 0.2490 | +0.44% |

Table.2-RL Efficiency Comparison: Fixed-Length vs. Gzip

| CR | Fixed | Gzip (standard) |  | Gzip (bitpacked) |  |
| --- | --- | --- | --- | --- | --- |
|  | BPS | BPS | Gain | BPS | Gain |
| 88 | 0.1250 | 0.1459 | -16.71% | 0.1240 | +0.82% |
| 58 | 0.1876 | 0.2213 | -17.97% | 0.1875 | +0.03% |
| 44 | 0.2501 | 0.2976 | -19.02% | 0.2491 | +0.38% |

Table.3-RL Entropy limitation Analysis

| CR | Fixed | Entropy Limit (H0) |  | Arithmetic Coder |  |
| --- | --- | --- | --- | --- | --- |
|  | BPS | BPS | Gain | BPS | Gain |
| 88 | 0.1250 | 0.1171 | +6.36% | 0.1218 | +2.56% |
| 58 | 0.1876 | 0.1802 | +3.95% | 0.1839 | +1.97% |
| 44 | 0.2501 | 0.2427 | +2.96% | 0.2477 | +0.96% |

Response to Review #2

- What makes the proposed codebook-projection method innovative? How does it differ from conventional vector quantization enhancements?

Response: The innovation of our Codebook Projection lies in its ability to perform learnable, adaptive feature selection prior to quantization. Conventional VQ methods often treat all latent dimensions equally, wasting codebook capacity on noisy features. Our approach explicitly projects embeddings into a subspace optimized for the codebook.

As demonstrated in our dimension masking experiment in Section V-B, Fig. 11 (d) on page 11, masking the vast majority of dimensions resulted in negligible performance variation, and strikingly, masking certain specific dimensions actually improved performance (manifesting as negative degradation). This observation suggests that the original high-dimensional embedding space is highly sparse and unrefined: most dimensions are redundant, while some even contain detrimental noise that obstructs effective quantization.

Consequently, relying on the full raw embedding space is inefficient. It is essential to filter out these redundant and harmful components to extract only the informative dimensions driving the reconstruction. Driven by this insight, our codebook projection mechanism acts as an adaptive, learnable feature selector. It projects the raw embeddings into a compact, optimized subspace prior to quantization. This design fundamentally differs from conventional brute-force VQ enhancements, as it prioritizes the quality of the features entering the codebook rather than merely increasing the codebook's capacity.

- How does the periodicity-aware modeling improve upon previous research work to leverage ECG periodicity?

Response: While previous works leverage ECG periodicity through explicit preprocessing (e.g., R-peak detection and fixed-window segmentation), our method improves upon this by introducing an implicit inductive bias via the Snake activation function. The improvements are supported by our mechanism analysis and simulation experiment.

Specifically, snake activation fx=x+ 1αsin2(αx) allows the model to naturally "resonate" with the ECG rhythm. The learnable parameter α dynamically adapts to the instantaneous heart rate (Heart Rate Variability), avoiding the need for rigid cutting.

Furthermore, standard activations like ReLU struggles to represent high-frequency sharp peaks. The sin2 term in Snake acts as a frequency multiplier, generating high-order harmonics. This Spectral Extension mechanism enables the network to efficiently synthesize the broad-spectrum components required to reconstruct sharp R-peaks without the smoothing artifacts seen in ReLU-based methods.

For empirically validation, we employed a simplified Multilayer Perceptron (MLP) and conducted a grid search experiment (varying hidden sizes from 20 to 100 and iterations from 1000 to 5000) to evaluate the fitting capability of MLPs that were identical except for their activation functions (ReLU vs. Snake) on an ECG signal. As shown by the dashed grey line in Fig. 1-RL, the ReLU network exhibits severe spectral bias. Regardless of model capacity or training duration, it results in an overly smoothed, near-linear approximation, completely failing to capture the sharp QRS complexes or the periodic rhythm. In strong contrast, the Snake network (blue lines) demonstrates distinct morphological fidelity. Even with limited capacity, it captures the oscillatory nature of the signal. With sufficient training (e.g., Iter=5000), it achieves precise reconstruction of the sharp peak amplitudes. This visual evidence confirms that the periodic inductive bias of Snake is essential for representing ECG signals efficiently, preventing the "smoothing" artifacts common in ReLU-based methods. The corresponding content has been added to Section V-C and Fig.12, page 12.

Fig.1-RL Comparative visualization of fitting performance across varying hidden sizes (Rows: 20, 50, 100) and iteration steps (Cols: 1k, 2k, 5k). The Orange line represents the ground truth ECG signal with high-frequency QRS complexes. The dashed Grey line (ReLU) exhibits spectral bias, resulting in overly smoothed approximations that fail to capture peak amplitudes. In contrast, the Blue line (Snake) successfully captures the oscillatory nature and sharp morphological details, demonstrating superior inductive bias for periodic biomedical signals.

- JR Cox, FM Nolle, HA Fozzard, and GC Oliver. Aztec, a preprocessing program for real-time ecg rhythm analysis. Biomedical Engineering, (2):128–129, 1968.

- Sateh MS Jalaleddine, Chriswell G Hutchens, Robert D Strattan, and William A Coberly. Ecg data compression techniques-a unified ap proach. IEEE transactions on Biomedical Engineering, 37(4):329–343, 1990.

- Peng He, Shaoming Meng, Yaping Cui, Dapeng Wu, and Ruyan Wang. Compression and encryption of heterogeneous signals for internet of medical things. IEEE Journal of Biomedical and Health Informatics, 28(5):2524–2535, 2024.

- Anukul Pandey. Ecg data compression using the formation of qrs complex segment bank and integer dct-based plateau region processing. Biomedical Signal Processing and Control, 85:104823, 2023.

- Ling Xiao, Qian Zhang, Kun Xie, and Chunxia Xiao. Online mecg compression based on incremental tensor decomposition for wear able devices. IEEE Journal of Biomedical and Health Informatics, 25(4):1041–1051, 2021.

- J. Gailly and M. Adler. zlib: A massively spiffy yet delicately unobtrusive compression library. GitHub repository, Version 1.3.1, 2023.

- P. Deutsch. RFC 1951 - DEFLATE Compressed Data Format Specification version 1.3. Technical report, Internet Engineering Task Force (IETF), 1996.

- GNU Project. gzip: The gzip compression utility. GNU Software, Version 1.12, 2023.

- P. Deutsch. RFC 1952 - GZIP file format specification version 4.3. Technical report, Internet Engineering Task Force (IETF), 1996.